Lessons Learned from Implementation of an Interferon Gamma Release Assay to Screen for Latent Tuberculosis Infection in a Large Multicenter Observational Cohort Study in Brazil

Abstract

Methods: This implementation study was performed within the RePORT-Brazil consortium between 2016 and 2019 across five research centers. The study described the operational challenges and solutions developed during the transition from QFT-GIT to the QuantiFERON-TB Gold Plus (QFT-Plus) assay. Two collection setups were compared: Setup A (same-site collection and processing) and Setup B (different locations requiring vehicle transport).

Results: Setup B centers experienced a higher percentage of indeterminate results and significantly more nonconformities, such as temperature deviations and longer transport times, compared to Setup A ($P < 0.001$). Temperature deviation was identified as the main nonconformity, particularly during spring and summer quarters. Over time, indeterminate results and nonconformities tended to decrease as the site personnel gained technical expertise.

Conclusion: Successful IGRA implementation requires rigorous standardization of preanalytical steps, regular technical training, and constant monitoring of transport time and temperature. These lessons facilitate large-scale LTBI screening in high TB-burden countries.

Keywords: tuberculosis, IGRA, QuantiFERON-Plus, LTBI, screening, quality control.

Lições Aprendidas com a Implementação de um Ensaio de Liberação de Interferon Gama para Triagem de Infecção Latente por Tuberculose em um Grande Estudo de Coorte Observacional Multicêntrico no Brasil

Resumo:

Métodos: Este estudo de implementação foi realizado no âmbito do consórcio RePORT-Brasil entre 2016 e 2019 em cinco centros de pesquisa. O estudo descreveu os desafios operacionais e as soluções desenvolvidas durante a transição do QFT-GIT para o ensaio QuantiFERON-TB Gold Plus (QFT-Plus). Foram comparadas duas configurações de coleta: Configuração A (coleta e processamento no mesmo local) e Configuração B (locais diferentes exigindo transporte por veículo).Resultados: Centros da Configuração B apresentaram um maior percentual de resultados indeterminados e significativamente mais não conformidades, como desvios de temperatura e tempos de transporte mais longos, em comparação com a Configuração A ($P < 0.001$). O desvio de temperatura foi identificado como a principal não conformidade, particularmente durante os trimestres de primavera e verão. Com o tempo, os resultados indeterminados e as não conformidades tenderam a diminuir à medida que o pessoal dos centros adquiriu experiência técnica.Conclusão: A implementação bem-sucedida do IGRA exige padronização rigorosa das etapas pré-analíticas, treinamento técnico regular e monitoramento constante do tempo e da temperatura de transporte. Essas lições facilitam a triagem de ILTB em larga escala em países com alta carga de TB.

Palavras-chave: tuberculose, IGRA, QuantiFERON-Plus, ILTB, triagem, controle de qualidade.